Digital Color Doppler Palm Ultrasound System

Model

SonoEye P5/V5/G5

USER MANUAL

Direction: CHUM SonoEye -001a

March.23rd ,2020

Rev. 1.0

CHISON Medical Technologies Co., Ltd.

We reserve the right to make changes to this manual without prior notice.

Regulatory Requirement

This product conforms to the essential requirements of the Medical Device Directive 93/42/EEC. Accessories without the CE mark are not guaranteed to meet the Essential Requirements of the Medical Device Directive.

This manual is a reference for the SonoEye P5/V5/G5. Please verify that you are using the latest revision of this document. If you need to know the latest revision, contact your distributor.

# TABLE OF CONTENTS

Chapter 1	Introduction	1

1.1. System Overview	1

1.2.Contact Information	1

Chapter 2  System Safety	2

2.1.	Safety Overview	2

2.2.	Electrical Safety	3

2.3.	Labels	5

2.3.1 Warning Symbols	5

2.3.2 Other Device Labels	7

2.4.	Patient Environmental Devices	7

2.5.	Biological Safety	9

2.6.	Scanning Patients and Education	10

Safe Scanning Guidelines	10

Understanding the MI/TI Display	12

Chapter 3  Preparing the System for Use	16

3.1.	Site Requirement	16

3.1.1.	Operation Environmental Requirement	16

3.1.2.	Electrical Requirements	16

3.2.	Downloading and Installing the App	17

3.3.	System Specifications	19

3.3.1.	Console Overview	19

3.3.2.	Physical Specifications	20

3.3.3.	Image Modes	20

3.3.4.	Accessories	20

3.3.5.	Configuration of the System	21

3.4.	System Positioning & Transporting	22

3.5.	Powering the System	23

3.5.1.	Acclimation Time	23

3.4.2  Connecting the electric power	23

3.6.	The system	23

Chapter 4  Control panel	25

4.1 Overview of display area	25

4.2 Create New Patient	26

4.3 Thumbnail Image	27

4.4	Probe	27

4.5 Switching Function	28

4.6 Switching Mode	28

4.7 Operating area of active state	28

4.8 Image parameter area	30

4.9 Cine memory bar	30

4.10 Soft iconboard	31

Chapter 5  Imaging	32

5.1	Select display mode	32

5.1.1 B mode	32

5.1.2 CFM Mode	32

5.1.3 PW mode	32

5.2 Select Examination mode	33

5.2.1 System identify	33

5.2.2 Select examination mode	34

5.3 Image control and adjustment	34

5.3.1 Gain	34

5.3.2	Depth	34

5.3.3 Frequency	34

5.3.4 Focus position	34

5.3.5 Zoom	34

5.3.6 TGC	35

5.3.6 B-Lines	35

5.3.6 Biopsy	37

5.4	Adjustment in PW mode	38

5.4.1 D Gain	38

5.4.3 Voice	38

5.4.4 Angle	38

5.4.5 Sampling Gate	38

5.4.6 Base line	38

Chapte 6 Measurement and Calculation	39

6.1	Measurement methods:	39

6.1.1Distance	39

6.1.2 Ellipse	39

6.2	B Mode measurement	39

6.2.1 Distance	39

6.2.1	Area	39

6.3	M Mode measurement	40

6.3.1	Distance	40

6.3.2	Time	40

6.3.3	HR	40

6.4	PW Mode measurement	40

6.4.1	Velocity	40

6.4.2	Time	40

Chapter 7 Cine-Memory	41

7.1	Store the image	41

7.2	Manual playback	41

7.3	Automatic playback	41

7.4	Cine Save/Recall	41

Chapter 8 Side Menu	42

8.1	Archive	42

8.2	EasyView	44

8.3	Setting	46

8.3.1Language setting	47

8.3.2Physical Keys Setting	48

8.3.3Doctor Accounts Management	49

8.4	Tutorials	49

8.5	Demo	50

8.6	About	50

Chapter 9  System Maintenance	51

9.1	Cleaning	51

9.2	System maintenance	51

9.3	Safety check	53

9.4	Troubleshooting	53

9.5	Service Responsibility	54

Chapter 10  The systems	55

10.1	General Description	55

10.2	Care and Maintenance	55

10.2.1	Inspecting Systems	55

10.2.2	Cleaning and Disinfecting	56

10.3	System Operation Instructions	62

10.3.1	Scanning the Patient	62

10.4	Service Responsibility	63

Appendix A: The Information of CE Representative	64

Appendix B: ACOUSTIC OUTPUT REPORT TABLE	65

Appendix C: Transducer Maximum Surface Temperature	69

Appendix D: MEASUREMENT RESULTS SUMMARY	70

Appendix E: GUIDANCE AND MANUFACTURER’S DECLARATION	71

# Chapter 1	Introduction

This manual contains necessary information for safe system operation.

Read and understand all instructions in this manual before operating the system. Always keep this manual with the equipment, and periodically review the procedures for operation and safety precautions.

## 1.1. System Overview

Indications for Use

SonoEye P5/V5/G5 Digital Color Doppler Palm Ultrasound System is intended for diagnostic ultrasound imaging in B (2D), Color Doppler, Combined(B+Color), Pulsed Wave and Tissue Harmonic Imaging modes. It is indicated for abdomen.

SonoEye P5/V5/G5 Digital Color Doppler Palm Ultrasound System is intended for use in environments where healthcare is provided by healthcare professionals.

Contraindication

The system is NOT intended for Ophthalmic use or any use that causes the acoustic beam to pass through the eye.

## 1.2.Contact Information

For additional information or assistance, please contact your local distributor or the appropriate support resource shown below:

| CHISON website | www.chison.com |
| --- | --- |
| Service Support | CHISON Medical Technologies Co., Ltd.. Tel: 0086-400-8878-020; 0086-0510-85311707 Fax: 0086-0510-85310726 E-mail: service@chison.com.cn |
| Placing an Order | CHISON Medical Technologies Co., Ltd. Tel: 0086-0510-8531-0593/0937 Fax: 0086-0510-85310726 Email: export@chison.com.cn |
| Manufacturer | CHISON Medical Technologies Co., Ltd. No.228, Changjiang East Road, Block 51 and 53, Phase 5, Shuofang Industrial Park, Xinwu District, Wuxi, Jiangsu, China 214142 No.9, Xinhuihuan Road, Xinwu District, Wuxi, Jiangsu, China 214028 |
| US Agent | Mr. Marco Mu, 2219 Rimland Drive,Suite 301,Bellingham, Barkley Villiage Bellingham,Washington,98226, UNITED STATES Phone: 360-3257028, Fax: 360-9253199, Email: us.agent@mid-link.net MID-LINK INTERNATIONAL CO., LTD |

Caution: Federal law restricts the device to sale by or on the order of a licensed practitioner or therapist.

# Chapter 2  System Safety

## Safety Overview

This section discusses measures to ensure the safety of both the operator and patient. To ensure the safety of both operator and patient, please read the relevant details in this chapter carefully before operating this system. Disregarding the warnings or violation of relevant rules may result in personal injury or even loss of life for operator or patient.

Users should observe the following precautions:

- This system complies with Type BF general equipment, and the IEC standard. Please follow Chapter 1 “System Safety” in the user’s manual to use this system properly.

- Do not modify this system in any way. The system is prohibiting to dismount. Necessary modifications must be made only by the manufacturer or its designated agents.

- This system has been fully adjusted at the factory. Do not adjust any fixed adjustable parts.

- In the event of a malfunction, turn off the system immediately and inform the manufacturer or its designated agents.

- Only connect this system, either electronically or mechanically, with devices that comply with the EN60601-1 standard. Recheck the leakage current and other safety performance indices of the entire system to avoid potential system damage caused by leakage from a current superposition.

- The system does not incorporate any specialized protective measures in the event it is configured with high-frequency operation devices. The operator should use caution in these types of applications.

- The system should be installed only by personnel authorized by the manufacturer. Do not attempt to install the system by yourself.

- Only an authorized service engineer may perform maintenance.

- Only a qualified operator, or someone under qualified supervision, should use the system.

- Do not use this system in the presence of flammable substances, otherwise an explosion may occur.

- Do not continuously scan the same part of a patient or expose the patient to prolonged scanning, otherwise it may harm the patient.

- When using the system for ultrasound testing, use only qualified ultrasound gel that complies with system standards.

- Do not unplug system when the system is in active operation. Always go to EXAM screen when need to remove the system.

- To prevent from arm or neck injury, the operator should not stay at the same position for too long during patient scanning without taking break.

- Do not put the system near the liquid.

NOTE

*The system has built-in screen saver to avoid the tic mark on the display. It is not recommended to constantly turn on and off the system.

*To dispose of this product properly, please call your local service department.

## Electrical Safety

Type of protection against electric shock

- Class I Equipment

CLASS I EQUIPMENT in which protection against electric shock does not rely on BASIC INSULATION only, but includes a protective earth ground. This additional safety precaution prevents exposed metal parts from becoming LIVE in the event of an insulation failure.

Degree of protection against electric shock

- Type BF Applied part (for Systems marked with BF symbol)

TYPE BF APPLIED PART providing a specified degree of protection against electric shock, with particular regard to allowable LEAKAGE CURRENT

BF: Isolation from ground; max. Patient leakage current: normal mode ≤100 µA, single fault condition ≤ 500 µA

Level of protection against harmful ingress of water

- The IP Classification of System is Ordinary Equipment , Immersed part is  IPX7, the other part is IPX1.

Safety level when used in the presence of FLAMMABLE ANAESTHETIC MIXED WITH AIR (or WITH OXYGEN or WITH NITROUS OXIDE):

The Equipment is not suitable for use in the environment with FLAMMABLE ANAESTHETIC MIXED WITH AIR (or WITH OXYGEN or WITH NITROUS OXIDE)

Mode of operation

- Continuous Operation

For maximum safety, always follow these guidelines:

- Do not remove or circumvent the grounding wire.

- Do not remove the protective covers on the system. These covers protect users from hazardous voltages. Cabinet panels must remain in place while the system is in use. A qualified electronic technician must make all internal replacements.

- Do not operate this system in the presence of flammable gases or anesthetics.

- All peripheral devices (unless certified as medical grade) that are connected to the system must be powered through the electrical outlet through an optional isolation transformer.

Notice upon Installation of Product

Separation distance and effect from fixed radio communications equipment: field strengths from fixed transmitters, such as base stations for radio (cellular/cordless) telephones and land mobile radios, amateur radio, AM and FM radio broadcast, and TV broadcast transmitter cannot be predicted theoretically with accuracy. To assess the electromagnetic environment due to fixed RF transmitters, an electromagnetic site survey should be considered. If the measured field strength in the location in which the ultrasound system is used exceeds the applicable RF compliance level as stated in the immunity declaration, the ultrasound system should be observed to verify normal operation. If abnormal operation is observed, additional measures may be necessary, such as re-orienting or relocating the ultrasound system or using an RF shielded examination room may be necessary.

- Use either power supply cords provided by or designated by CHISON. Products equipped with a power source plug should be plugged into the fixed power socket which has the protective grounding conductor. Never use any adaptor or converter to connect with a power source plug (e.g. three-prong-to-two-prong converter).

- Locate the equipment as far away as possible from other electronic equipment.

- Be sure to use only the cables provided by or designated by CHISON. Connect these cables following the installation procedures (e.g. wire power cables separately from signal cables).

- Lay out the main equipment and other peripherals following the installation procedures described in this manual.

Notice against User Modification

The user should never modify this product.

User modifications may cause degradation in Electrical Safety. Modification of the product includes changes in:

- Cables (length, material, wiring, etc.)

- System configuration/components

User modifications may cause degradation in EMC performance. Modification of the product includes changes in:

- Cables (length, material, wiring, etc.)

- System installation/layout

- System configuration/components

- Securing system parts (cover open/close, cover screwing)

## Labels

### 2.3.1 Warning Symbols

| Icon | Meaning |
| --- | --- |
|  | Refer to instruction manual/booklet. |
|  | Caution, consult accompanying documents. This symbol advises the reader to consult the accompanying documents for important safety related information such as warnings and pre-cautions that cannot be presented on the device itself. |
|  | The CE mark of Conformity indicates this equipment conforms to the Council Directive 93/42/EEC. |
|  | Serial number of the device. |
|  | This symbol is followed by the manufacturing date of the device in the form YYYY-MM. |
|  | This symbol is accompanied by the name and the address of the manufacturer. |

### 2.3.2 Other Device Labels

The following table describes the purpose and location of safety labels and other important information provided on the equipment.

Table 2-1: Label Icons

| Icon | Meaning |
| --- | --- |
| Identification and Rating Plate | • Manufacture’s name • Serial number |
|  | Type-BF applied part |
| IPX7 | Immersed part is IPX7 |

## Patient Environmental Devices

Profile side :

  - System

  - Power switch

Acceptable Devices

The Patient Environmental devices shown above are specified to be suitable for use within the PATIENT ENVIRONMENT.

CAUTION:

- DO NOT connect any systems or accessories without approval by CHISON within the PATIENT ENVIRONMENT.

- DO NOT touch patient and devices without IEC/EN 60601-1 approval to avoid the leakage current risk within the PATIENT ENVIRONMENT.

Unapproved Devices

CAUTION:

- DO NOT use unapproved devices.

- If devices are connected without the approval of CHISON, the warranty will be INVALID.

- The system can’t be used with HF surgical equipment, otherwise the burns to patient may occur.

Any device connected to this system must conform to one or more of the requirements listed below:

- IEC standard or equivalent standards appropriate to devices.

CAUTION: Unsafe operation or malfunction may result. Use only the accessories, options and supplies approved or recommended in these instructions for use.

Peripheral used in the patient environment

The system has been verified for overall safety, compatibility and compliance with the following on-board image recording devices:

- B/W video printer:  Mitsubishi P93W; Sony UP-897MD

- Color video printer:  Mitsubishi CP31W (CP30W)

The system may also be used safely while connected to devices other than those recommended above if the devices and their specifications, installation, and interconnection with the system conform to the requirements of IEC/EN 60601-1-1.

The connection of equipment or transmission networks other than as specified in the user instructions can result in an electric shock hazard or equipment malfunction. Substitute or alternate equipment and connections require verification of compatibility and conformity to IEC/EN 60601-1-1 by the installer. Equipment modifications and possible resulting malfunctions and electromagnetic interference are the responsibility of the owner.

General precautions for installing an alternate off-board, remote device or a network would include:

- The added device(s) must have appropriate safety standard conformance and CE Marking.

- There must be adequate mechanical mounting of the device and stability of the combination.

- Risk and leakage current of the combination must comply with IEC/EN 60601-1.

- Electromagnetic emissions and immunity of the combination must conform to IEC/EN 60601-1-2.

Peripheral used in the non-patient environment

The system has been verified for compatibility, and compliance for connection to a local area network (LAN) via a wire LAN, provided the LAN components are IEC/EN 60950 compliant.

General precautions for installing an alternate off-board, remote device or a network would include:

- The added device(s) must have appropriate safety standard conformance and CE Marking.

- The added device(s) must be used for their intended purpose having a compatible interface.

## Biological Safety

This product, as with all diagnostic ultrasound equipment, should be used only for valid reasons and should be used both for the shortest period of time and at the lowest power settings necessary (ALARA - As Low As Reasonably Achievable) to produce diagnostically acceptable images. The AIUM offers the following guidelines:

Clinical Safety Quoted from AIUM

Approved March 26, 1997

Diagnostic ultrasound has been in use since the late 1950s. Given its known benefits and recognized efficacy for medical diagnosis, including use during human pregnancy, the American Institute of Ultrasound in Medicine herein addresses the clinical safety of such use:

There are no confirmed biological effects on patients or instrument operators caused by exposures from present diagnostic ultrasound instruments. Although the possibility exists that such biological effects may be identified in the future, current data indicate  that the benefits to patients of the prudent use of  diagnostic ultrasound  outweigh  the risks, if  any that may be present.

Heating: Elevating tissue temperature during obstetrical examinations creates medical concerns.  At the embryo development stage, the rise in temperature and the length of time exposed to heat combine to determine potential detrimental effects. Exercise caution particularly during Doppler/Color exams. The Thermal Index (TI) provides a statistical estimate of the potential temperature elevation (in centigrade) of tissue temperature. Three forms of TI are available: Soft Tissue Thermal Index (TIS), Bone Thermal Index (TIB).

Soft Tissue Thermal Index (TIS). Used when imaging soft tissue only, it provides an estimate of potential temperature increase in soft tissue.

Bone Thermal Index (TIB). Used when bone is near the focus of the image as in the third trimester OB examination, it provides an estimate of potential temperature increase in the bone or adjacent soft tissue.

Cavitation: Cavitation may occur when sound passes through an area that contains a cavity, such as a gas bubble or air pocket (in the lung or intestine, for example). During the process of cavitation, the sound wave may cause the bubble to contract or resonate. This oscillation may cause the bubbles to explode and damage the tissue. The Mechanical Index (MI) has been created to help users accurately evaluate the likelihood of cavitation and the related adverse effects.

MI recognizes the importance of non-thermal processes, cavitation in particular, and the Index is an attempt to indicate the probability that they might occur within the tissue.

## Scanning Patients and Education

The Track-3 or IEC60601-2-37 output display standard allows users to share the responsibility for the safe use of this ultrasound system. Follow these usage guidelines for safe operation:

- In order to maintain proper cleanliness of the systems, always clean them between patients.

- Always use a disinfected sheath on all EV/ER systems during every exam.

- Continuously move the system, rather than staying in a single spot, to avoid elevated temperatures in one part of the patient’s body.

- Move system away from the patient when not actively scanning.

- Understand the meaning of the TI, TIS, TIB and MI output display, as well as the relationship between these parameters and the thermal/cavitation bioeffect to the tissue.

- Expose the patient to only the very lowest practical transmit power levels for the shortest possible time to achieve a satisfactory diagnosis (ALARA - As Low As Reasonably Achievable).

### Safe Scanning Guidelines

- Ultrasound should only be used for medical diagnosis and only by trained medical personnel.

- Diagnostic ultrasound procedures should be done only by personnel fully trained in the use of the equipment, in the interpretation of the results and images, and in the safe use of ultrasound (including education as to potential hazards).

- Operators should understand the likely influence of the machine controls, the operating mode (e.g. B-mode, color Doppler imaging or spectral Doppler) and system frequency on thermal and cavitation hazards.

- Select a low setting for each new patient. Output should only be increased during the examination if penetration is still required to achieve a satisfactory result, and after the Gain control has been moved to its maximum value.

- Maintain the shortest examination time necessary to produce a useful diagnostic result.

- Do not hold the system in a fixed position for any longer than is necessary. It should be removed from the patient whenever there is no need for real-time imaging or spectral Doppler acquisition. The frozen frame and Cine loop capabilities allow images to be reviewed and discussed without exposing the patient to continuous scanning.

- Take particular care to reduce output and minimize exposure time of an embryo or fetus when the temperature of the mother is already elevated.

- Take particular care to reduce the risk of thermal hazard during diagnostic ultrasound when exposing: an embryo less than eight weeks after gestation; or the head, brain or spine of any fetus or neonate.

- Operators should continually monitor the on-screen thermal index (TI) and mechanical index (MI) values and use control settings that keep these settings as low as possible while still achieving diagnostically useful results. In obstetric examinations, TIS (soft tissue thermal index) should be monitored during scans carried out in the first eight weeks after gestation, and TIB (bone thermal index) thereafter.

MI>  0.3	There is a possibility of minor damage to neonatal lung or intestine. If such exposure is necessary, reduce the exposure time as much as possible.

MI>  0.7	There is a risk of cavitation if an ultrasound contrast agent containing gas micro-spheres is being used. There is a theoretical risk of cavitation without the presence of ultrasound contrast agents. The risk increases with MI values above this threshold.

TI> 0.7	The overall exposure time of an embryo or fetus should be restricted in accordance with Table 2-2 below as a reference:

| TI | Maximum exposure time (minutes) |
| --- | --- |
| 0.7 | 60 |
| 1.0 | 30 |
| 1.5 | 15 |
| 2.0 | 4 |
| 2.5 | 1 |

Table 2-2 Maximum recommended exposure times for an embryo or fetus

- Non-diagnostic use of ultrasound equipment is not generally recommended. Examples of non-diagnostic uses of ultrasound equipment include repeated scans for operator training, equipment demonstration using normal subjects, and the production of souvenir pictures or videos of a fetus. For equipment of which the safety indices are displayed over their full range of values, the TI should always be less than 0.5 and the MI should always be less than 0.3. Avoid frequent repeated exposure of any subject. Scans in the first trimester of pregnancy should not be carried out for the sole purpose of producing souvenir videos or photographs, nor should their production involve increasing the exposure levels or extending the scan times beyond those needed for clinical purposes.

- Diagnostic ultrasound has the potential for both false positive and false negative results. Misdiagnosis is far more dangerous than any effect that might result from the ultrasound exposure. Therefore, diagnostic ultrasound system should be performed only by those with sufficient training and education.

### Understanding the MI/TI Display

Track-3 follows the Output Display Standard for systems that include fetal Doppler applications. The acoustic output will not be evaluated on an application-specific basis, but the global maximum de-rated Ispta must be ≤ 720 mW/cm2 and either the global maximum MI must be ≤ 1.9 or the global maximum de-rated Ispta must be ≤ 190 W/cm2. An exception is for ophthalmic use, in which case the TI = max (TIS_as) is not to exceed 1.0; Ispta.3 ≤50mW/cm2, and MI ≤ 0.23. Track-3 gives the user the freedom to increase the output acoustic power for a specific exam, and still limit output acoustic power within the global maximum de-rated Ispta ≤ 720 mW/cm2 under an Output Display Standard.

For any diagnostic ultrasonic systems, Track-3 provides an Output Indices Display Standard. The diagnostic ultrasound systems and its operator’s manual contain the information regarding an ALARA (As Low As Reasonably Achievable) education program for the clinical end-user and the acoustic output indices, MI and TI. The MI describes the likelihood of cavitation, and the TI offers the predicted maximum temperature rise in tissue as a result of the diagnostic examination. In general, a temperature increase of 2.5°C must be present consistently at one spot for 2 hours to cause fetal abnormalities. Avoiding a local temperature rise above 1°C should ensure that no thermally induced biologic effect occurs. When referring to the TI for potential thermal effect, a TI equal to 1 does not mean the temperature will rise 1 degree C. It only means an increased potential for thermal effects can be expected as the TI increases. A high index does not mean that bioeffects are occurring, but only that the potential exists and there is no consideration in the TI for the scan duration, so minimizing the overall scan time will reduce the potential for effects. These operator control and display features shift the safety responsibility from the manufacturer to the user. So it is very important to have the Ultrasound systems display the acoustic output indices correctly and the education of the user to interpret the value appropriately.

RF: (De-rating factor)

In Situ intensity and pressure cannot currently be measured. Therefore, the acoustic power measurement is normally done in the water tank, and when soft tissue replaces water along the ultrasound path, a decrease in intensity is expected. The fractional reduction in intensity caused by attenuation is denoted by the de-rating factor (RF),

RF  = 10 (-0.1 a f z)

Where a is the attenuation coefficient in dB cm-1 MHz-1, f is the transducer center frequency, and z is the distance along the beam axis between the source and the point of interest.

De-rating factor RF for the various distances and frequencies with attenuation coefficient 0.3dB cm-1 MHz-1 in homogeneous soft tissue is listed in the following table. An example is if the user uses 7.5MHz frequency, the power will be attenuated by .0750 at 5cm, or 0.3x7.5x5=-11.25dB. The De- rated Intensity is also referred to as ‘.3’ at the end (e.g. Ispta.3).

Distance	Frequency (MHz)

(cm) 	1 	3 	5 	7.5

| 1 | 0.9332 | 0.8128 | 0.7080 | 0.5957 |
| --- | --- | --- | --- | --- |
| 2 | 0.8710 | 0.6607 | 0.5012 | 0.3548 |
| 3 | 0.8128 | 0.5370 | 0.3548 | 0.2113 |
| 4 | 0.7586 | 0.4365 | 0.2512 | 0.1259 |
| 5 | 0.7080 | 0.3548 | 0.1778 | 0.0750 |
| 6 | 0.6607 | 0.2884 | 0.1259 | 0.0447 |
| 7 | 0.6166 | 0.2344 | 0.0891 | 0.0266 |
| 8 | 0.5754 | 0.1903 | 0.0631 | 0.0158 |

I’=I*RF Where I’ is the intensity in soft tissue, I is the time-averaged intensity measured in water.

Tissue Model:

Tissue temperature elevation depends on power, tissue type, beam width, and scanning mode. Six models are developed to mimic possible clinical situations.

Thermal Models  Composition            Mode 	Specification 	Application

| 1 | TIS | Soft tissue | Unscanned | Large aperture (>1cm2) | Liver PW |
| --- | --- | --- | --- | --- | --- |
| 2 | TIS | Soft tissue | Unscanned | Small aperture (<1cm2) | Pencil System |
| 3 | TIS | Soft tissue | Scanned | Evaluated at surface | Breast color |
| 4 | TIB | Soft tissue and bone | Scanned | Soft tissue at surface | Muscle color |
| 5 | TIB | Soft tissue and bone | Unscanned | Bone at focus | Fetus head PW |

Soft tissue:

Describes low fat content tissue that does not contain calcifications or large gas-filled spaces.

Scanned: (auto-scan)

Refers to the steering of successive burst through the field of view, e.g. B and color mode.

Unscanned:

Emission of ultrasonic pulses occurs along a single line of sight and is unchanged until the transducer is moved to a new position. For instance, the PW, CW and M mode.

TI:

TI is defined as the ratio of the In Situ acoustic power (W.3) to the acoustic power required to raise tissue temperature by 1°C (Wdeg), TI = W.3/Wdeg.

Three TIs corresponding to soft tissue (TIS) for abdominal; bone (TIB) for fetal and neonatal cephalic; have been developed for applications in different exams.

An estimate of the acoustic power in milliwatts necessary to produce a 1°C temperature elevation in soft tissue is:

Wdeg  = 210/fc,	for model 1 to 4, where fc is the center frequency in MHz.

Wdeg  = 40 K D	for model 5 and 6, where K (beam shape factor) is 1.0, D is the aperture diameter in cm at the depth of interest.

MI:

Cavitation is more likely to occur at high pressures and low frequencies in pulse ultrasound wave in the tissue, which contains the bubble or air pocket (for instance, the lung, intestine, or scan with gas contrast agents). The threshold under optimum conditions of pulsed ultrasound is predicted by the ration of the peak pressure to the square root of the frequency.

MI = Pr’ / sqrt(fc)

Pr’ is the de-rated (0.3) peak rare-fractional pressure in Mpa at the point where PII is the maximum, and fc is the center frequency in MHz. PII is the Pulse Intensity Integral that the total energy per unit area carried by the wave during the time duration of the pulse. The peak rare- fractional pressure is measured in hydrophone maximum negative voltage normalized by the hydrophone calibration parameter.

Display Guideline:

For different operation modes, different indices must be displayed. However, only one index needs to be shown at a time.  Display is not required if maximum MI is less than 1.0 for any setting of the operating mode, or if maximum TI is less than 1.0 for any setting of the operating mode. For TI, if the TIS and TIB are both greater than 1.0, the scanners need not be capable of displaying both indices simultaneously. If the index falls below 0.4, no display is needed.

Display and Report in Different Mode

Located on the upper middle section of the system display monitor, the acoustic output display provides the operator with real-time indication of acoustic levels being generated by the system.

For B-Scan Mode

Only display and report MI, and start from 0.4 if maximum MI > 1.0, display in increments of 0.2.

For Color Mode

Only display and report TIS or TIB and start from 0.4 if maximum TI > 1.0, display in increments of 0.2 for values of indices of 2.0 or less, and 0.5 for values of indices greater than 2.0.

For Doppler Mode

Only display and report TIS or TIB and start from 0.4 if maximum TI > 1.0, display in increments of 0.2 for values of indices of 2.0 or less, and 0.5 for values of indices greater than 2.0.

Below is a simple guideline for the user when TI exceeds one limit exposure time to 4(6-TI) minutes based on the ‘National Council on Radiation Protection. Exposure Criteria for Medical Diagnostic Ultrasound: I. Criteria Based on Thermal Mechanisms. Report No.113 1992’.

Operator Control Features:

The user should be aware that certain operator controls may affect the acoustic output. It is recommended to use the default (or lowest) output power setting and compensate using Gain control to acquire an image. Other than the output power setting in the soft-menu, which has the most direct impact on the power; the PRF, image sector size, frame rate, depth, and focal position also slightly affect the output power. The default setting is normally around 70% of the allowable power depending on the exam application mode.

Controls Affecting Acoustic Output

The potential for producing mechanical bioeffects (MI) or thermal bioeffects (TI) can be influnced by certain controls.

Direct: The Acoustic Output control has the most significant effect on Acoustic Output.

Indirect: Indirect effects may occur when adjusting controls. Controls that can influence MI and TI are detailed under the Bioeffects portion of each control in the Optimizing the Image chapter.

Always observe the Acoustic Output display for possible effects.

Best practices while scanning

HINTS: Raise the Acoustic Output only after attempting image optimization with controls that have no effect on Acoustic Output, such as Gain and TGC.

WARNING: Be sure to have read and understood control explanations for each mode used before attempting to adjust the Acoustic Output control or any control that can effect Acoustic Output.

Use the minimum necessary acoustic output to get the best diagnostic image or measurement during an examination. Begin the exam with the system that provides an optimum focal depth and penetration.

Acoustic Output Default Levels

In order to assure that an exam does not start at a high output level, the system initiates scanning at a reduced default output level. This reduced level is preset programmable and depends upon the exam icon and system selected. It takes effect when the system is powered on or New Patient is selected. To modify acoustic output, adjust the Power Output level on the Soft Menu.

# Chapter 3  Preparing the System for Use

## Site Requirement

### Operation Environmental Requirement

The following environmental conditions are within system tolerances for operation:

Strong radiation sources or powerful electromagnetic waves (e.g. electro-magnetic waves from radio broadcasting) may result in image ghosting or noise. The system should be isolated from such radiation sources or electromagnetic waves.

|  | Environment | Operation | Saving and Transporting |
| --- | --- | --- | --- |
| Parameter |  |  |  |
| Temperature | 10℃～40℃ | -10℃～50℃ |
| Relative Humidity | 30%～75% | ≤80%，non-condensing |
| Atmosphere Pressure | 700hPa～1060hPa | 700hPa～1060hPa |

### Electrical Requirements

Power Requirements：DC 5V

Power Consumption： ≤10VA

Voltage Fluctuation：±5%

WARNING:

Maintain a fluctuation range of less than ±10% of voltage labeling on rear panel of the system, otherwise the system may be damaged.

Built in battery specifications:

| Battery model | BT-100 |
| --- | --- |
| Capacity | 4100mAh |
| Rated voltage | 3.8V |
| Standard charge voltage | 4.35V |
| Discharge closing voltage | 3V |
| Discharge time | About 150 minutes |
| Standard charge current | 4A |
| Maximum continuous discharge current | 4A |
| Battery structure | 1P |
| Package shell | PC |
| Cycle life | 500 times |
| Charging time | About 120±20 minutes |
| Operating temperature | Charge | 0℃～45℃；32oF～113oF |
|  | Discharge | -10℃～60℃；-4oF～149oF |
|  | Storage | -20℃～45℃；-4oF～113oF |
| Battery status indicator | Light up 3 lights | Power balance : 75%~100% |
|  | Light up 2 lights | Power balance : 40%~75% |
|  | Light up 2 lights | Power balance : 0%~40% |
|  | Turn off 3 lights | Power balance : <0% |
|  | Charging tips | Charge indicator |

Adapter specifications:

| Adapter model | Mascot 3626-50 |
| --- | --- |
| Input | 100-240V～Max.0.6A 50/60Hz |
| Output | 15V4A |

## Downloading and Installing the App

Downloading the App

Use a computer to obtain the.apk installation package from chison and move it to a mobile terminal

Installing the App

- Select the file which ended by .apk to install APP.

- After installation, press the icon to enter the interface.

Giving App Access to Shared Device Storage

App uses shared device storage for the patient database and to access your device's camera for QR code scanning.

Some versions of the Android operating system require that you specify that an app is allowed access to shared device storage. If your device prompts you to allow App to access photos, media, or files on your device, touch Allow. If you touch Deny, you cannot use App until you give access to shared device storage in the Android App Permissions settings.

Updating the App

Use a computer to obtain the.apk update package from chison and move it to the mobile terminal. Select the update package to install APP.

Turning the System On and Off

- Power on:

Press  for 3 seconds to start system.

- Power off:

Press  for 3 seconds to shut down the system.

## System Specifications

### Console Overview

Fig. 3-1 a: Console Overview

Fig. 3-1: Firmware Front View

From left to right: Freeze Button, customize button: it can be set as Freeze, Gain, Depth and save Image

### Physical Specifications

Dimensions of main unit (approx.): 60mm（Width）*172mm（Length）*24mm（Height）

### Image Modes

- B mode

- M mode

- Color Doppler Imaging 

- Power Doppler Imaging

### Accessories

Transducers:

P5     2.0MHz- 4.5MHz     Convex Array

V5     2.0MHz- 4.5MHz     Convex Array

G5     2.0MHz- 4.5MHz     Convex Array

### Configuration of the System

| Function | SonoEye P5 | SonoEye V5 | SonoEye G5 |
| --- | --- | --- | --- |
| B mode | Standard | Standard | Standard |
| M mode | Option | Standard | Option |
| PW mode | Option | Standard | Option |
| CFM mode | Option | Option | Option |
| FHI | Option | Option | Standard |
| Type-C cable | Standard | Standard | Standard |
| Gain | Standard | Standard | Standard |
| Depth | Standard | Standard | Standard |
| TGC | Standard | Standard | Standard |
| Focus | Standard | Standard | Standard |
| Compound | Standard | Standard | Standard |
| Frequency Scaling | Standard | Standard | Standard |
| Freeze | Standard | Standard | Standard |
| Zoom | Standard | Standard | Standard |
| Maximum display depth(≥19.5) | Standard | Standard | Standard |
| B-distance | Standard | Standard | Standard |
| B-Circumference | Standard | Standard | Standard |
| B-Area | Standard | Standard | Standard |
| B-Volume | Standard | Standard | Standard |
| M-Distance | Option | Option | Option |
| M-Time | Option | Option | Option |
| M-HR | Option | Option | Option |
| PW-Velocity | Option | Option | Option |
| PW-Time | Option | Option | Option |
| Vessel measure package | Option | Option | Option |
| Thyroid measure package | Option | Option | Option |
| Abdomen measure package | Option | Option | Option |
| Cardiac measure package | Option | Option | Option |
| Lung measurement | Option | Option | Option |
| General report | Standard | Standard | Standard |
| Vessel report | Option | Option | Option |
| Abdomen report | Option | Option | Option |
| Cadiac report | Option | Option | Option |
| Lung report | Option | Option | Option |
| Chinese Interface | Standard | Standard | Standard |
| Multi-language Interface | Option | Option | Option |
| Clipboard | Standard | Standard | Standard |
| Auto IMT | Option | Option | Option |
| Instant AIO | Option | Option | Option |
| Biopsy Guide Line | Option | Option | Option |
| One-click full display | Option | Option | Option |
| Save Cine | Standard | Standard | Standard |
| Save Image | Standard | Standard | Standard |
| Cine Loop | Option | Option | Option |
| Intuitive instant touch adjust STC | Option | Option | Option |
| Intuitive instant adjust Depth | Option | Option | Option |
| Intuitive instant adjust focus position | Option | Option | Option |
| Intuitive instant bi directional Cine loop | Option | Option | Option |
| Intuitive instant touch zoom | Option | Option | Option |
| Intuitive instant adjust cine loop (manually/ Automatically) | Option | Option | Option |
| Intuitive instant adjust sample Gate location | Option | Option | Option |
| Intuitive instant adjust size of sample Gate horizontal | Option | Option | Option |
| Intuitive instant adjust size of sample Gate vertical | Option | Option | Option |
| Intuitive instant adjust PW sample gate | Option | Option | Option |

## System Positioning & Transporting

Moving the System

When moving or transporting the system, take the precautions described below to ensure maximum safety for personnel, the system and other equipment.

Before Moving the System

- Completely switch off the system.

NOTE

To prevent damage to the power cord, DO NOT pull excessively on the cord or sharply bend the cord while wrapping it.

  - Store all systems in their original cases or wrap them in soft cloth or foam to prevent damage.

  - Replace gel and other essential accessories in the appropriate storage case.

  - Ensure that no loose items are left.

When Moving the System

- Take extra care when you move the system long distances and on inclines. Ask for help if necessary.

Avoid ramps that are too steeper to tip over the system. Utilize additional care and personnel when moving on steep incline (> 5 º ) or loading into a vehicle for transport.

- Use extra care when crossing door or elevator thresholds.

CAUTION

  - Walk slowly and carefully when moving the system.

  - Be sure the pathway is clear.

  - Do not let the system strike walls or doorframe.

Transporting the System

After preparing the system as described above, take the following additional precautions:

- Before transporting, place the system in its original storage carton.

- While the system is on a lift, ensure that the lift is capable of bear the weight of the system and the passengers. To avoid the movement of the system, secure the system by using wood chocks, restraining straps or other similar types of constraints.

- Drive carefully to prevent damage from vibration. Avoid unpaved roads, excessive speeds, and erratic stops or starts.

## Powering the System

### Acclimation Time

After being transported, the unit requires one hour for each 2.5 ºC increment if its temperature is below 10 ºC or above 40 ºC.

### 3.4.2  Connecting the electric power

After making sure the power supply on displayer is normal status, and the voltage type is matched to the power requirement indicated on the label of system, then connect the connector with displayer.

## The system

CAUTION

Before connecting the system, please carefully check the system lens, system cable and system connector to see whether there is anything abnormal, such as cracks, falls off. Abnormal system is not allowed to connect to the system; otherwise there is possibility of electricity shock.

- Connect the connector from the P5/V5/G5 system to displayer.

CAUTION

- Only power supply at “turn off” state, can install / take-down the system system, otherwise it will damage the machine or the system.

- If system is not correctly or completely inserted to the system, or if system is not securely connected to the socket, this may cause mis-operation, e.g. the system cannot be recognized by the system, mis-recognized, or the system may drop off from the main unit and be damaged.

- When installing and disassembling system, please put the system head inside the system holder, it can prevent the system falling down to the ground.

Deactivating the System

Disconnect the connector from displayer to deactivating the system.

# Chapter 4  Control panel

## 4.1 Overview of display area

Fig 4.1a Display Interface

| Side Menu | 2. Probe | 3. Image Functional Area |
| --- | --- | --- |
| Full | Center Line | Save Image |
| Save Cine | Set the Exam Mode | Freeze |
| Logo | Patient ID | End Exam |
| Connection Icon | Parameter Display Area | Thumbnail image |
| Focus position |  |  |

## 4.2 Create New Patient

Press  to end the current exam, And start a new patient.

To add patient information:

- On the imaging display, touch patient ID.

- On the Patient Info display, type the patient information.

- Click the [OK] button.

Patient Information:

| ID | Patient ID |
| --- | --- |
| Date | Current date |
| Fist Name | Input patient’s First Name |
| Middle Name | Input patient’s Middle Name |
| Last Name | Input patient’s Middle Name |
| Sex | Select the patient’s sex |
| Age | Set the patient’s age, the system will automatically calculate the patient’s birthday |
| Birth | Set the patient’s birthday, the system will automatically calculate the patient’s age |
| Height | Input the patient’s height |
| Weight | Input the patient’s weight |

## 4.3 Thumbnail Image

After saving the image or cine, press  to check the saved image and cine.

## Probe

Press  to display all probe information. Touch an exam preset you want.

## 4.5 Switching Function

Slide the top left along the buttons to show Freq.

## 4.6 Switching Mode

Available imaging modes are B, C, and D-Mode.

Press B, then select the corresponding mode.

## 4.7 Operating area of active state

| Button | Name | Function |
| --- | --- | --- |
|  | TGC | Swipe from right to left to show the TGC button, move the slider left or right to adjust TGC. |
|  | Freeze | Touch this icon to freeze the current image |
|  | Freq. | Press this button to adjust the Freq.(4.0M/6.5M/7.5M/8.5M/10M/FHI) |
|  | B Mode | Press this key to select the exam mode. |
|  | C Mode | Press C to enter the CFM Mode. |
|  | M Mode | Press M to enter the M Mode. |
|  | D Mode | Press D to enter PW Mode. |
|  | Side Menu | Press to enter the side menu list. |
|  | Full | Press the key to enlarge the exam area. |
|  | Measure Button | Press this key to enter measurement. |
|  | Ellipse | Press the Ellipse to start Ellipse measurement. |
|  | Distance | Press the Dist to start distance measurement. |
|  | Delete | Press Del. to delete all the measurement. |
|  | Mark | Press this button and the user can select the content to add in current exam interface. |
|  | Comment | Press this button to add comment in current exam interface. |
|  | Arrow | Press this button to add arrow in current exam interface. |
|  | BodyMark | Press this button to add body mark in current exam interface. |
|  | Delete | Press this button to clear all the content that user have added in in current exam interface. |

## 4.8 Image parameter area

Display different parameters under different mode. The different display mode has different parameters.

## 4.9 Cine memory bar

Display the frame of current cine memory. By clicking on the screen manually or automatically to playback or touch play control icon.

## 4.10 Soft iconboard

Fig 4.1b Soft iconboard

| CapsLock | Digital character iconboard switching icon | Chinese and English switching icon |
| --- | --- | --- |
| The Backspace icon | The Searching icon |

Caution

The system has Chinese input defaults to the double-spelling input.

# Chapter 5  Imaging

This chapter will introduce image display modes and the operation of image control and adjustment.

## Select display mode

There are five image display modes: B, M, color mode, FHI and PW mode, can be shifted by the mode icon.

### 5.1.1 B mode

Click [B] icon to display single B mode image. B mode is the basic operating mode for two-dimensional scanning and diagnosis.

### 5.1.2 CFM Mode

Color Flow Map is a technique for imaging blood flow by displaying flow data such as velocity and direction on B mode image. Based on Doppler Effect, normally the blood flow moving toward the probe scan direction is marked in red, while blood flow moving away from probe scan direction is marked in blue. Touch [C] icon, screen only displays color mode operation interface.

Fig. 5-1 CFM Mode

### 5.1.3 PW mode

Intended Use:

Doppler is intended to provide measurement data concerning the velocity of moving tissues and fluids. PW Doppler lets you examine blood flow data selectively from a small region called the Sample Volume.

The X axis represents time while the Y axis represents velocity in either a forward or reverse direction.

PW Doppler is typically used for displaying the speed, direction, and spectral content of blood flow at selected anatomical sites.

PW Doppler can be combined with B-mode for quickly selecting the anatomical site for PW Doppler examination. The site where PW Doppler data is derived appears graphically on the B-mode image (Sample Volume Gate). The Sample Volume Gate can be moved anywhere within B-mode image.

PW mode Exam Procedure:

- Connect the appropriate system, leaving the systems in their respective holders.

- Position the patient for the examination.

- Press PATIENT-ID and enter the appropriate patient data.

- Select the application and system to be used.

- Locate the anatomy to be examined. Get a good B-mode image. Press C- icon to help locate the vessel you wish to examine.

- Press D- icon to display the sample volume cursor and gate.

- Position or re-size the sample volume gate by moving the slide left, right, up and down.

- Press D- icon to display PW Doppler spectrum and the system operates in combined B+Doppler mode. The Doppler signal is heard through the speakers.

- Optimize the PW Doppler spectrum, as necessary.

- Press the corresponding button to transfer the exam mode between real time B-mode with Doppler mode (with audio).

- Sample along the whole length of the vessel. Ensure that the system is parallel to flow. Listen, then look, when positioning the sample volume cursor.

- Perform measurements and calculations, as necessary.

- Record results with your recording devices.

- Press  to resume imaging.

- Repeat the above procedure until all relevant flow sites have been examined.

- Replace the system in its respective holder.

### 5.2 Select Examination mode

### 5.2.1 System identify

When you connect one system, the app will identify the system automatically.

Caution

When inserting the system, please make sure the surface which has CHISON logo is always upside.

### 5.2.2 Select examination mode

Click icon ,the examination interface will pop-up, according to the type of probe and the sex of people, different examination will be activated .Click the icon to enter into the preset which you want.

### 5.3 Image control and adjustment

### 5.3.1 Gain

To adjust the Gain:

Swiping left, right in arbitrary place of screen.

Press the Gain button and slide the block to adjust the Gain.

### Depth

To adjust the Depth:

Swiping Dynamic range is used for adjusting the contrast resolution of B mode image and mode image, compressing or enlarging the display range of gray scale.

At the real-time status, drug the slide of dynamic to adjust the gain value from 30 to 90 dB.16 steps for adjust.

### 5.3.3 Frequency

Click  to choose Frequency. The range of the frequency depends on different probe.

### 5.3.4 Focus position

Click on focus icon and slide the block to change the focus position

Press and hold the target position of the ruler, the focus will automatically jump to that position or click [Focus], drag the slider to adjust the focus position in exam area.

### 5.3.5 Zoom

- Choose the step by clicking the icon .

- Touch the screen with two fingers and slide it outward to adjust the zoom position.

### 5.3.6 TGC

Fig 1-7  8TGC

TGC curves can be used for adjusting gain compensation in different image dept.

Side the TGC block to adjust it

### 5.3.6 B-Lines

Select the Lung examination image in Easyview, the B-Lines, Pleura-Line, Diffuse and Delete button will appear at the bottom of image.

Press B-Lines button, the B-Lines and line space will display on screen, and the right bottom of image will display pleura data.

Press Pleura-Line button, the pleura location will display on screen, and the right bottom of image will display pleura data.

Press Diffuse button, the screen displays below picture and the right bottom of image display gray scale histogram.

After finish all the measurement, the report button will display on the bottom of picture. Press it, the system will display current patient’s report.

### 5.3.6 Biopsy

Press menu to enter the biopsy, the biopsy guide line appeared, press biopsy button again to exist Biopsy.

Adjust guide line:

Press the button to edit, slide it to adjust the location.

Super Needle: Super needle is used for enhance the needle image in the B mode image. After turn on the super needle, super needle and needle angle function will be active and user can adjust the needle angle to optimize the image for needle only

## Adjustment in PW mode

### 5.4.1 D Gain

At real-time state, slide left or right in spectrum area to adjust the size of color gain, adjustment range from 0~255, the smallest unit of adjustment is 1.

### 5.4.3 Voice

At real-time state, press the Voice button on displayer to adjust the voice.

### 5.4.4 Angle

At real-time state, press the adjust icon on the left of the screen to adjust the angle.

### 5.4.5 Sampling Gate

At real-time state, use two finger to slide it outward to adjust the size of sampling gate.

At real-time state, press the sampling gate and move it to adjust the location.

### 5.4.6 Base line

At real-time state, press the base line and move it to top and bottom to adjust the location.

# Chapte 6 Measurement and Calculation

Main content of this chapter:

Measurement in B Mode.

## Measurement methods:

The system contains Distance, Ellipse.

### 6.1.1Distance

Measurement steps:

- Click the『Distance』icon under the measurement menu to enter into measurement.

- Click the B image area; it will display a segment with two “+” icon. One of the “+” is active, you can move it by dragging your finger to fit the one point of the line.

- After measurement, the result will appear on the exam interface.

- Repeat the above step to start a new measurement, press  can delete the current measurement result.

- Press  can delete all the measurement.

### 6.1.2 Ellipse

Measurement steps:

- Click the [Ellipse] icon under the measure menu to enter into measurement.

- Click in the B image area，it will display a ellipse with 4 “+”，you can move the “+” by dragging your finger on B image area to fit it’s position.

- After measurement, the result will appear on the exam interface.

- Repeat the above step to start a new measurement, press  can delete the current measurement result.

- Press can delete all the measurement.

## B Mode measurement

Press B to enter the B mode and press measure to start measurement.

### 6.2.1 Distance

Distance measurement is the same as section 6.1.1

### Area

Area measurement is the same as section 6.1.2

## M Mode measurement

### Distance

This feature allows the measurement of the distance between two points. It is a measurement between the two horizontal lines that lean on the two cursors. The position of the vertical time line does not affect the distance measurement.

Distance measurement is the same as section 6.1.1

### Time

Time is the measurement between the two vertical time lines created by two cursors. The position of the horizontal distance line does not affect time measurements.

### HR

HR is the measurement between the two vertical lines that are created by two cursors in beat per minute (BPM). The position of the horizontal distance line does not affect HR.

## PW Mode measurement

### Velocity

Press Freeze after finish the scan, select the velocity in menu, the blue “+” displayed on the screen, move it to desired positon to start measurement. The result will display on image area and result area. Repeat the above steps to start new measurement.

### Time

Press Freeze after finish the scan, select the time in menu, there are 2 blue “+” displayed on the screen, move it to desired positon to start measurement. The result will display on image area and result area. Repeat the above steps to start new measurement.

# Chapter 7 Cine-Memory

This chapter introduces the theory of saving images in Cine-memory and the operation of image playback in Cine- memory.

## Store the image

Images in B-mode can be stored in Cine-memory at the unit of frame in time sequence. If the storage is full of images, when storing a latest new frame image, the first saved frame image will be removed out of Cine-memory. Therefore there are always the latest images in the storage. All the images in Cine-memory can be played back manually or automatically.

Fig 7-1 Movie playback bar diagram

## Manual playback

After clicking the  button to freeze the image, the movie playback bar pops up. At this time, you can click and drag in the image area. Swipe right to display the images in ascending order of frames, that is, the same order as the images are stored, otherwise Frames are displayed in descending order.

## Automatic playback

Press Freeze and click  to start automatic playback.

## Cine Save/Recall

Press Freeze and click  to saving current cine.

# Chapter 8 Side Menu

The setting function is used to running environment, state and configuration of each exam mode while system is start up. The system setting was saved in memory and it still exist when system is power off.

## Archive

Press  on the top left corner of the screen, select archive to enter the archive management.

Press  on the top right corner of the screen, user can manage the patient list according to ID, MRN and Name.

Select Patient: Press the Patient ID and hold it, the current patient was selected.

Select all patient: Press can select all patient.

Delete Patient: Press  can delete the current patient.

Press Patient ID to enter EasyView.

## EasyView

Press EasyView to check the patient information and exam image.

Press the top right corner of the displayer, the user can set the image preview mode.

Press the image and hold it to select the current image, the frame of image will turn to blue.

: Press this button to send current image to external media device.

: Press this button to print current image.

: Press this button to select all image.

: Press this button to delete current image.

## Setting

Press the setting button to enter the setting interface.

### 8.3.1Language setting

Press language to select the needed language.

### 8.3.2Physical Keys Setting

Select the physical keys setting to set the key function.

| Button | Option |
| --- | --- |
| “+” | “Not set”“Freeze”“Gain+”“Depth+”“Depth-”“Save Image” |
| “-” |  |
| “P” |  |

### 8.3.3Doctor Accounts Management

Click the Doctor Accounts Management to manage the accounts.

## Tutorials

The first time you start the app, it displays tutorials to familiarize you with the features of the system.You can view the Tutorials at any time. Touch , and then click Tutorials to enter the interface as below:

## Demo

Click Demo to enter the interface, the user can view the exam image and cine.

Exit Demo: press the anywhere on screen, then press “<” on the top left corner.

## About

Press About to check the current software Version.

.

# Chapter 9  System Maintenance

## Cleaning

Caution

Before clean any part of the system, please make sure that the system is turned off and the power cord is disconnected from the power supply socket. Otherwise there will be danger of electricity shock.

Cleaning method of the iconboard, outer-frame, system holder and monitor:

Please use a piece of soft and dry cloth to clean the system. If there’s some dirty difficult to be cleaned, please use wet cloth to clean system, and then use dry cloth to wipe off the water on the system.

Caution

Please don’t use organic solvent to clean the system; otherwise it will damage the system surface.

Please never allow any liquid get inside the system or system, otherwise it will damage the system and cause electronic short.

If the system connector, STC slide or any peripheral device is required to be cleaned, please contact our authorized agent in your country in advance. Any cleaning by unauthorized person may result in system malfunction or affect its features.

## System maintenance

According to the purposes, the systems of this system can be divided into 2 categories: those used on the surface of patient body, or those used inside the body of patient.

Caution

No matter which type of examination is performed, please always try to reduce the unnecessary radiation of ultrasound wave to the patient during the ultrasound examination

Caution

- System can only be used by professional doctor who has received professional training of ultrasound.

- It is forbidden to sterilize and disinfect system by high pressure. If it needs to be used in sterilized occasion, please use a sterilized disposable system cover on the system.

- Please avoid drop off or hitting the system by anything.

- Don’t scratch the system surface while using it.

- Please use the authorized ultrasound gel during scanning. Using un-authorized gel may cause scratch or damage to system surface.

- Please keep the system clean and dry.

- Please don’t use or preserve the system where it is over 50℃.

- Please carefully check the system surface, before using. If there is any abnormal phenomenon (eg. there’s a leakage on the system surface), please stop using the system immediately and contact our authorized agent in your country as soon as possible. If you don’t know the contact number of your authorized agent, please contact us by detail contact information at the end of this chapter.

System maintenance

Please take good care of the system. Collision and dropping is strongly prohibited.

Please use the ultrasound gel which is acknowledged by the manufacture of the unit. We recommend AQUASONIC Gel made by R. P. Kincheloe Company in USA.

Plug and unplug of system in real-time is strongly prohibited.

Wash the system:

1) System tip

Rinsing: rinse the surface with running water, and use a sponge or soft cloth to remove

gently the dirt and gel on system tip

2) Connector, Cable, other part of the system tip must not be soaked in a solution. Simply clean it using a soft cloth moistened with alcohol and then dry it.

Disinfection: When necessary, soaking the system tip in disinfection solution. The recommended disinfection solution is CIDEX ACTIVATED DIALDEHYDE SOLUTION. (Manufacturer: Johnson and Johnson Medical). It has been approved by the FDA, its 510(K) number is K924434. The following instructions are provided by Johnson and Johnson Medical. For more detail, please contact Johnson and Johnson Medical

1) Soaking temperature: 10(C~40(C

2) Atmospheric pressure: 700hPa~ 1060hPa

3) Soaking time: Under FDA requirements, CIDEX Activated Dialdehyde Solution requires 45 minutes processing at 25(C for high level disinfection. The 45 minutes processing time was established as the time to kill 6 logs of the test organism.

Rinsing: Sufficiently rinse the system by using water to remove chemicals.

Aeration and let the system become dry in normal temperature.

Please strictly keep the system away from the paint thinner, ethylene oxide, other organic solvent, etc

Please keep the system inside the system case when it is not in use.

Dipping the system into any liquid is strongly prohibited.

Caution

Please immediately stop using the system and system if there is any broken phenomenon on the electricity cable or the system transducer. Otherwise there will be a danger of the electricity shock.

## Safety check

To ensure the system work normally, please make a maintenance plan, check the safety of the system periodically. If there is any abnormal phenomenon with the machine, please contact our authorized agent in your country as soon as possible.

If there is no image or menu on the screen or other phenomenon appears after switching on the machine, please do troubleshooting first according to the following check list. If the trouble is still not solved, please contact our authorized agent in your country as soon as possible.

## Troubleshooting

According to the most frequently occurred errors and system messages, the list of possible causes and relevant solutions is attached as below:

| Errors & Messages | Possible Cause | Solution |
| --- | --- | --- |
| When turn on the system, power-indicating lamp is not lit. | Type c cord may not be connected, or may not be well connected to the displayer. | Please contact with Sales office, service department and distributor |
| When turn on the system, power indicating lamp is lit, but no images on the monitor. | The system is restarted too shortly after it is switched off. | Wait for 1 minute to restart the system after switching off. |
| Menu bar displays on the screen but no scanning image | Transmission frequency, gain or STC control is not set properly. System is connected improperly the system is in frozen status | Adjust the transmission frequency, gain or STC control. Ensure the system is connecting correctly Defreeze the system by pressing the FREEZE icon. |
| Image quality is abnormal | Examination mode is not correct Image processing parameters are set improperly | Set the exam mode correctly. Adjust the setting of image processing or set it to default setting |
| Wrong probe type show on screen. | Internal circuit protection | Restart machine |

## Service Responsibility

If users install, use and maintain the system fully according to CHISON’s installation manual, operation manual and service manual, then SonoEye P5/V5/G5 main unit has a life time of 5 years,

The warranty of the system after ex-work is as the time in the warranty card.

The system is a precise electronic system. Only the CHISON’s authorized service engineer could replace the defective parts. Any assembly, disassembly, handling, repair, or replacement by any other people may have adverse impact on the safety and effectiveness of the systems and transducers, and thus will reduce the life time of the system and transducers, and such systems and transducers will not be covered by CHISON warranty after the above improper handling. Standard maintenance must be performed by CHISON’s authorized service engineer during the life time of the product.

# Chapter 10  The systems

## General Description

Fig.6-1: Convex Probe Overview

The probes provide high spatial and contrast ultrasound imaging of frequencies from 2.0MHz to 4.5MHz. These systems operate by pulsing sound waves into the body and listening to the returning echoes to produce high-resolution brightness mode, and a real time display.

## Care and Maintenance

The systems that come with the system are designed to be durable and dependable. These precision instruments should be inspected daily and handled with care. Please observe the following precautions:

- Do not drop the transducer on hard surface. This can damage the transducer elements and compromise the  electrical safety of the transducer.

- Avoid kinking or pinching the transducer cable.

- Use only approved ultrasonic coupling gels.

- Follow the instructions for cleaning and disinfecting that come with each system.

### Inspecting Systems

Before and after each use, inspect carefully the system’s lens, cable, casing, and connector. Look for any damage that would allow liquid to enter the system. If any damage is suspected, do not use the system until it has been inspected and repaired/replaced by an authorized Service Representative.

NOTE

Keep a log of all system maintenance, along with a picture of any system malfunction.

WARNING

The systems are designed to be used only with this ultrasound system. Use of these systems on any other system or a non-qualified system may cause electrical shock or damage on the system/transducer.

### Cleaning and Disinfecting

- Place the system into the solution of cleaning-disinfectant. Make sure not to immerse the system into the liquid beyond the immersion level given in the pictures below. Make sure that the system is covered with the cleaning-disinfectant up to the immersion level during the complete disinfection time.

- For the recommended cleaning and disinfection time, please see your Operating Manual.

- Scrub the system as needed using a soft sponge, gauze, or cloth to remove all visible residue from the system surface.

- Rinse the system with enough clean, potable water to remove all disinfectant residues.

- Use a soft cloth to clean the cable and the user section of the system with the cleaning disinfectant liquid. Make sure that the surface of the system and cable is wetted thoroughly with the cleaning-disinfectant.

- Allow system to air dry completely.

- Reconnect the system to the ultrasound console and place the system into its holder.

System Immersion Levels

Recommended materials for cleaning and disinfecting transducer

| Solution | Origin | Qualified Use | Active Ingredients | Purpose | Time |
| --- | --- | --- | --- | --- | --- |
| WIP'ANIOS | FRA | Wipe | Isopropyl alcohol/Didecyl dimethyl ammonium chlorides | Disinfection | NA |
| 80% Ethanol（Using clean cloth soak the liquid） | Any | Wipe | Ethanol | Clean or Disinfection | NA |
| 70% Isopropyl alcohol（Using clean cloth soak the liquid） | Any | Wipe | Isopropyl alcohol | Clean or Disinfection | NA |
| Clinell Sporicidal Wipes | GBR | Wipe | Sodium Percarbonate Citric Acid | Disinfection | NA |
| mikrozid PAA wipes | DEU | Wipe | peracetic acid,Hydrogen peroxid,Acetic acid | Disinfection | NA |
| Universal wipes | GBR | Wipe | Quaternary ammonium | Disinfection | NA |
| perform classic wipes EP | DEU | Wipe | EP | Disinfection | NA |
| MetriCide OPA Plus | USA | Soak | Orthophthalal-aldehyde | Disinfection | 12 |
| CIDEZYME | Any | Soak | enzyme | Clean | 1 |
| Perform classic concentrate GA | DEU | Soak | glutaraldehyde | Disinfection | 30 |
| CIDEX OPA | Any | Soak | Orthophthalal-aldehyde | Disinfection | 12 |
| Hydrogen peroxid | Sultan Healthcare | Soak | Hydrogen peroxid | Disinfection | 30 |
| gigasept FF(new) | DEU | Soak | Succin diald ehyde | Disinfection | 15 |
| neodisher mediclean | DEU | Soak | enzyme | Clean | 10 |
| ANIOSYME XL3 | FRA | Soak | enzyme | Clean | 5 |
| Epizyme Rapid(Rapid Multi-Enzyme cleaner） | Any | Soak | enzyme | Clean | 10 |
| MetriZyme | USA | Soak | enzyme | Clean | 5 |
| Endozime Xtreme Power | USA | Soak | enzyme | Clean | 2 |

CAUTION

These transducers are not designed to withstand heat sterilization methods. Exposure to temperatures in excess of 60 º C will cause permanent damage. The transducers are not designed to be totally submerged in fluid, as permanent damage will result if the entire transducer is submerged.

System Safety

Handling precautions

Ultrasound systems are highly sensitive medical instruments that can easily be damaged by improper handling. Use care when handling and protect from damage when not in use. DO NOT use a damaged or defective system. Failure to follow these precautions can result in serious injury and equipment damage.

Electrical shock hazard:

The system is driven with electrical energy that can injure the patient or user if live internal parts are contacted by conductive solution:

- DO NOT immerse the system into any liquid beyond the level indicated by the immersion level diagram. Never immerse the system connector into any liquid.

- Prior to each use, visually inspect the system lens and case area for cracks, cuts, tears, and other signs of physical damage. DO NOT use a system that appears to be damaged until you verify functional and safe performance. You need to perform a more thorough inspection, including the cable, strain relief, and connector, each time you clean the system.

- Before inserting the connector into the system port, inspect the system connector pins. If a pin is bent, DO NOT use the system until it has been inspected and repaired/replaced by a CHISON Service Representative.

- Electrical leakage checks should be performed on a routine basis by CHISON Service or qualified hospital personnel.

Mechanical hazard:

A defective system or excess force can cause patient injury or system damage:

- Inspect systems for sharp edges or rough surfaces that may injure sensitive tissue.

- DO NOT apply excessive force to the system connector when inserting into the system port. The pin of a system connector may bend.

Special handling instructions

Using protective sheaths

The use of market cleared system sheaths is recommended for clinical applications. Reference FDA March 29, 1991 "Medical Alert on Latex Products".

Protective sheaths may be required to minimize disease transmission. System sheaths are available for use with all clinical situations where infection is a concern. Use of legally marketed, sterile system sheaths is strongly recommended for endo-cavitary procedures.

DO NOT use pre-lubricated condoms as a sheath. In some cases, they can damage the system. Lubricants in these condoms may not be compatible with system construction.

Devices containing latex may cause severe allergic reaction in latex sensitive individuals. Refer to FDA’s March 29, 1991 Medical Alert on latex products.

DO NOT use an expired system sheath. Before using a sheath, verify if it has expired.

System handling and infection control:

This information is intended to increase user awareness of the risks of disease transmission associated with using this equipment and provide guidance in making decisions directly affecting the safety of the patient as well as the equipment user.

Diagnostic ultrasound systems utilize ultrasound energy that must be coupled to the patient by direct physical contact.

Depending on the type of examination, this contact occurs with a variety of tissues ranging from intact skin in a routine exam to recirculating blood in a surgical procedure.  The level of risk of infection varies greatly with the type of contact.

One of the most effective ways to prevent transmission between patients is with single use or disposable devices.  However, ultrasound transducers are complex and expensive devices that must be reused between patients. It is very important, therefore, to minimize the risk of disease transmission by using barriers and through proper processing between patients.

Risk of Infection

ALWAYS clean and disinfect the system between patients to the level appropriate for the type of examination and use FDA-cleared system sheaths where appropriate.

Adequate cleaning and disinfection are necessary to prevent disease transmission. It is the responsibility of the equipment user to verify and maintain the effectiveness of the infection control procedures in use. Always use sterile, legally marketed system sheaths for intra-cavitary procedures.

System Cleaning process:

DO disconnect the system from the system prior to cleaning/disinfecting the system. Failure to do so could damage the system.

Perform Cleaning system after each use

- Disconnect the system from the ultrasound console and remove all coupling gel from the system by wiping with a soft cloth and rinsing with flowing water.

- Wash the system with mild soap in lukewarm water. Scrub the system as needed using a soft sponge, gauze, or cloth to remove all visible residue from the system surface. Prolonged soaking or scrubbing with a soft bristle brush (such as a toothbrush) may be necessary if material has dried onto the system surface.

WARNING

To avoid  electrical  shock,  always  turn  off  the  system  and  disconnect  the system before cleaning the system.

CAUTION

Take extra care when handling the lens face of the Ultrasound transducer. The lens face is especially sensitive and can easily be damaged by rough handling. NEVER use excessive force when cleaning the lens face.

- Rinse the system with enough clean potable water to remove all visible soap residue.

- Air dry or dry with a soft cloth.

CAUTION

To minimize the risk of infection from blood-borne pathogens, you must handle the system and all disposables that have contacted blood, other potentially infectious materials, mucous membranes, and non-intact skin in accordance with infection control procedures. You must wear protective gloves when handling potentially infectious material. Use a face shield and gown if there is a risk of splashing or splatter.

Disinfecting the systems:

After each use, please disinfect the systems. Ultrasound systems can be disinfected using liquid chemical germicides. The level of disinfection is directly related to the duration of contact with the germicide. Increased contact time produces a higher level of disinfection.

In order for liquid chemical germicides to be effective, all visible residue must be removed during the cleaning process. Thoroughly clean the system, as described earlier before attempting disinfection.

You MUST disconnect the system from the system prior to cleaning/disinfecting the system. Failure to do so could damage the system.

DO NOT soak systems in liquid chemical germicide for longer than is stated by the germicide instructions for use. Extended soaking may cause system damage and early failure of the enclosure, resulting in possible electric shock hazard.

- Prepare the germicide solution according to the manufacturer's instructions. Be sure to follow all precautions for storage, use and disposal. The transducer is not designed to be totally submerged in fluid. Permanent damage will result if the entire transducer is submerged. The immersed part shall not exceed the transducer binding line.

- Place the cleaned and dried system in contact with the germicide for the time specified by the germicide manufacturer.

- After removing from the germicide, rinse the system following the germicide manufacturer's rinsing instructions.  Flush all visible germicide residue from the system and allow to air dry.

Ultrasound transducers can easily be damaged by improper handling and by contact with certain chemicals.  Failure to follow these precautions can result in serious injury and equipment damage

- Do not immerse the system into any liquid beyond the level specified for that system.

- Avoid mechanical shock or impact to the transducer and do not apply excessive bending or pulling force to the cable.

- Transducer damage can result from contact with inappropriate coupling or cleaning agents:

- Do not soak or saturate transducers with solutions containing alcohol, bleach, ammonium chloride compounds or hydrogen peroxide

- Avoid contact with solutions or coupling gels containing mineral oil or lanolin

- Avoid temperatures above 60°C. Under no circumstances should the transducer be subjected to heat sterilization method. Exposure to temperatures above 60º C will cause permanent damage to the transducer.

- Inspect the system prior to use for damage or degeneration to the housing, strain relief, lens and seal.  Do not use a damaged or defective system.

Coupling gels

DO NOT use unrecommended gels (lubricants). They may damage the system and void the warranty. AQUASONIC Gel made by R. P. Kincheloe Company in USA is recommended.

In order to assure optimal transmission of energy between the patient and system, a conductive gel must be applied liberally to the patient where scanning will be performed.

DO NOT apply gel to the eyes. If there is gel contact to the eye, flush eye thoroughly with water.

Coupling gels should not contain the following ingredients as they are known to cause system damage:

- Methanol, ethanol, isopropanol, or any other alcohol-based product.

- Mineral oil

- Iodine

- Lotions

- Lanolin

- Aloe Vera

- Olive Oil

- Methyl or Ethyl Parabens (para hydroxybenzoic acid)

- Dimethylsilicone

Planned maintenance

The following maintenance plan is suggested for the system and systems to ensure optimum operation and safety.

Daily: inspect the systems

After each use: clean the systems, disinfect the systems.

As necessary: inspect the systems, clean the systems, disinfect the systems.

Returning/Shipping Systems and Repair Parts

Transportation dept. and our policy require that equipment returned for service MUST be clean and free of blood and other infectious substances.

When you return a system or part for service, you need to clean and disinfect the system or part prior to packing and shipping the equipment.

Ensure that you follow system cleaning and disinfection instructions provided in this Manual.

This ensures that employees in the transportation industry as well as the people who receive the package are protected from any risk.

## System Operation Instructions

For details on connecting, activating, deactivating, disconnecting, transporting and storing the systems, see Section 3.5 “Systems” in Chapter 3.

### Scanning the Patient

In order to assure optimal transmission of energy between the patient and system, a conductive gel must be applied liberally to the patient where scanning will be performed.

After the examination is complete, follow the cleaning and disinfecting, or sterilizing procedures as appropriate.

## Service Responsibility

The system is a precise electronic system. Only an authorized service contractor should replace defective parts. Failures caused by unauthorized service are not the responsibility of the manufacturer.

REFERENCE:

  - AIUM/NEMA:  Standard For Real-Time Display of Thermal and Mechanical Acoustic Output Indices On Diagnostic Ultrasound Equipment, Revision 2.  NEMA Standards Publication UD 3-2004; American Institute of Ultrasound in Medicine, Laurel MD; National Electrical Manufacturers Association, Rosslyn, VA; 2004a.

  - Implementation of the Principle of As Reasonably Achievable (ALARA) for Medical and Dental Personnel, National Council on Radiation Protection and Measurements (NCRP), report NO.107, December 31,1990

  - FDA Center for Devices and radiological Health (CDRH), 510(K) Guidance for Diagnostic Ultrasound and Fetal Doppler Ultrasound Medical Devices, September 8 1989 draft

  - FDA/CDRH,510(K) Diagnostic Ultrasound Guidance Update of 1991, April 26, 1991 draft

  - Biological Effects of Ultrasound: Mechanisms and Clinical Implications, NCRP Report No. 74, December 30,1983

  - Exposure Criteria for Medical Diagnostic Ultrasound: I. Criteria Based on Thermal Mechanisms, NCRP Report No.113, June 1,1992

  - Bioeffects Considerations for the safety of Diagnostic Ultrasound, Journal of Ultrasound in Medicine, AIUM, September1988

  - Geneva Report on Safety and Standardization in Medical Ultrasound, WFUMB, May 1990 Medical Ultrasound Safety, AIUM, 1994

  - Medical Electrical Equipment standard IEC 60601-1,  IEC60601-1-2, IEC 60601-2-37

  - Diagnostic Ultrasound Physics and Equipment, edit by P. R. Hoskins, in 2003

# Appendix A: The Information of CE Representative

Shanghai International Trading Corp. GmbH (Hamburg)

Add: Eiffestrasse 80, 20537  Hamburg，Germany

Tel: 0049-40-2513175

Fax:0049-40-255726

E-mail: antonsissi@hotmail.com  shholding@hotmail.com

Appendix B: ACOUSTIC OUTPUT REPORT TABLE

Transducer Model: P5/V5/G5

Operation Mode: B

| Index Label | MI | TIS | TIB | TIC |
| --- | --- | --- | --- | --- |
|  |  | At Surface | Below Surface | At Surface | Below Surface |  |
| Maximum Index Value | 1.05 | 0.11 | 0.11 | 0.25 |
| Index component Value |  | 0.11 | 0.11 | 0.11 | 0.11 |  |
| Acoustic Parameters | pr.α at ZMI | (MPa) | 1.58 |  |  |  |  |  |
|  | P | (mW) |  | 10.02 | 10.02 | 12.62 |
|  | P1*1 | (mW) |  | 10.02 | 10.02 |  |
|  | Zs | (cm) |  |  | 2.15 |  |  |  |
|  | Zb | (cm) |  |  |  |  | 2.15 |  |
|  | ZMI | (cm) | 2.15 |  |  |  |  |  |
|  | Zpii.α | (cm) | 2.15 |  |  |  |  |  |
|  | fawf | (MHz) | 2.25 | 2.25 | 2.25 | 2.22 |
| Other Information | prr | (Hz) | - |  |  |  |  |  |
|  | srr | (Hz) | 18.94 |  |  |  |  |  |
|  | npss | NA | 1.00 |  |  |  |  |  |
|  | Ipa.α at Zpii.α | (W/cm2) | 79.17 |  |  |  |  |  |
|  | Ispta.α at Zpii.α or Zsii.α | (mW/cm2) | 0.79 |  |  |  |  |  |
|  | Ispta at Zpii or Zsii | (mW/cm2) | 0.82 |  |  |  |  |  |
|  | pr at Zpii | (MPa) | 1.79 |  |  |  |  |  |
| Operating control conditions | Focus | (cm) | 3.00 | 3.00 | 3.00 | 3.00 |
|  | Depth | (cm) | 16.00 | 16.00 | 16.00 | 16.00 |
|  | Freq | MHz | 2.00 | 2.00 | 2.00 | 2.00 |
|  | PRF | HZ | - | - | - | - |

Transducer Model: P5/V5/G5

Operation Mode:B+ M

| Index Label | MI | TIS | TIB | TIC |
| --- | --- | --- | --- | --- |
|  |  | At Surface | Below Surface | At Surface | Below Surface |  |
| Maximum Index Value | 1.12 | 0.26 | 0.58 | 0.36 |
| Index component Value |  | 0.26 | 0.18 | 0.26 | 0.58 |  |
| Acoustic Parameters | pr.α at ZMI | (MPa) | 1.65 |  |  |  |  |  |
|  | P | (mW) |  | 20.12 | 20.12 | 20.12 |
|  | P1*1 | (mW) |  | 14.58 | 14.58 |  |
|  | Zs | (cm) |  |  | 1.52 |  |  |  |
|  | Zb | (cm) |  |  |  |  | 2.22 |  |
|  | ZMI | (cm) | 2.05 |  |  |  |  |  |
|  | Zpii.α | (cm) | 2.05 |  |  |  |  |  |
|  | fawf | (MHz) | 2.16 | 2.16 | 2.16 | 2.16 |
| Other Information | prr | (Hz) | 346 |  |  |  |  |  |
|  | srr | (Hz) | - |  |  |  |  |  |
|  | npss | NA | 1.00 |  |  |  |  |  |
|  | Ipa.α at Zpii.α | (W/cm2) | 80.46 |  |  |  |  |  |
|  | Ispta.α at Zpii.α or Zsii.α | (mW/cm2) | 8.77 |  |  |  |  |  |
|  | Ispta at Zpii or Zsii | (mW/cm2) | 8.89 |  |  |  |  |  |
|  | pr at Zpii | (MPa) | 1.75 |  |  |  |  |  |
| Operating control conditions | Focus | (cm) | 3.00 | 3.00 | 3.00 | 3.00 |
|  | Depth | (cm) | 16.00 | 16.00 | 16.00 | 16.00 |
|  | Freq | MHz | 2.00 | 2.00 | 2.00 | 2.00 |
|  | PRF | HZ | - | - | - | - |

Transducer Model: P5/V5/G5

Operation Mode: B+C

| Index Label | MI | TIS | TIB | TIC |
| --- | --- | --- | --- | --- |
|  |  | At Surface | Below Surface | At Surface | Below Surface |  |
| Maximum Index Value | 0.88 | 0.22 | 0.22 | 0.50 |
| Index component Value |  | 0.22 | 0.22 | 0.22 | 0.22 |  |
| Acoustic Parameters | pr.α at ZMI | (MPa) | 1.39 |  |  |  |  |  |
|  | P | (mW) |  | 24.21 | 24.21 | 28.90 |
|  | P1*1 | (mW) |  | 18.68 | 18.68 |  |
|  | Zs | (cm) |  |  | 3.20 |  |  |  |
|  | Zb | (cm) |  |  |  |  | 3.20 |  |
|  | ZMI | (cm) | 2.60 |  |  |  |  |  |
|  | Zpii.α | (cm) | 2.60 |  |  |  |  |  |
|  | fawf | (MHz) | 2.51 | 2.48 | 2.48 | 2.38 |
| Other Information | prr | (Hz) | - |  |  |  |  |  |
|  | srr | (Hz) | 6.49 |  |  |  |  |  |
|  | npss | NA | 8.00 |  |  |  |  |  |
|  | Ipa.α at Zpii.α | (W/cm2) | 94.79 |  |  |  |  |  |
|  | Ispta.α at Zpii.α or Zsii.α | (mW/cm2) | 0.50 |  |  |  |  |  |
|  | Ispta at Zpii or Zsii | (mW/cm2) | 0.51 |  |  |  |  |  |
|  | pr at Zpii | (MPa) | 1.73 |  |  |  |  |  |
| Operating control conditions | Focus | (cm) | 3.00 | 3.00 | 3.00 | 5.00 |
|  | Depth | (cm) | 16.00 | 16.00 | 16.00 | 16.00 |
|  | Freq | MHz | 3.00 | 3.00 | 3.00 | 3.00 |
|  | PRF | HZ | - | - | - | - |

Transducer Model: P5/V5/G5

Operation Mode: PW

| Index Label | MI | TIS | TIB | TIC |
| --- | --- | --- | --- | --- |
|  |  | At Surface | Below Surface | At Surface | Below Surface |  |
| Maximum Index Value | 0.63 | 0.20 | 0.74 | 0.38 |
| Index component Value |  | 0.15 | 0.20 | 0.16 | 0.74 |  |
| Acoustic Parameters | pr.α at ZMI | (MPa) | 1.00 |  |  |  |  |  |
|  | P | (mW) |  | 21.79 | 18.92 | 22.59 |
|  | P1*1 | (mW) |  | 12.61 | 14.60 |  |
|  | Zs | (cm) |  |  | 1.40 |  |  |  |
|  | Zb | (cm) |  |  |  |  | 3.20 |  |
|  | ZMI | (cm) | 3.05 |  |  |  |  |  |
|  | Zpii.α | (cm) | 3.05 |  |  |  |  |  |
|  | fawf | (MHz) | 2.48 | 2.47 | 2.50 | 2.45 |
| Other Information | prr | (Hz) | 2000 |  |  |  |  |  |
|  | srr | (Hz) | - |  |  |  |  |  |
|  | npss | NA | 1.00 |  |  |  |  |  |
|  | Ipa.α at Zpii.α | (W/cm2) | 40.29 |  |  |  |  |  |
|  | Ispta.α at Zpii.α or Zsii.α | (mW/cm2) | 109.67 |  |  |  |  |  |
|  | Ispta at Zpii or Zsii | (mW/cm2) | 213.87 |  |  |  |  |  |
|  | pr at Zpii | (MPa) | 1.22 |  |  |  |  |  |
| Operating control conditions | Focus | (cm) | 4.00 | 11.00 | 4.00 | 7.00 |
|  | Depth | (cm) | 16.00 | 16.00 | 16.00 | 16.00 |
|  | Freq | MHz | 3.00 | 3.00 | 3.00 | 3.00 |
|  | PRF | HZ | 2000 | 2000 | 2000 | 2000 |

Appendix C: Transducer Maximum Surface Temperature

| Transducer model | Maximum surface temperature(℃) Contacting human-tissue mimicking material | Maximum surface temperature(℃) Suspending in air |
| --- | --- | --- |
| P5 | <37.6 | <31.4 |
| V5 | <37.6 | <31.4 |
| G5 | <37.6 | <31.4 |

Appendix D: MEASUREMENT RESULTS SUMMARY

| Measurement | Useful Range | Accuracy |
| --- | --- | --- |
| Distance | Image area | <±5% |
| Circumference: trace method,ellipse method | Image area | <±5% |
| Area: trace method,ellipse method | Image area | <±10% |
| Volume | Image area | <±5% |
| Angle | Image area | <±5% |

Appendix E: GUIDANCE AND MANUFACTURER’S DECLARATION

| 1. Guidance and manufacturer’s declaration – electromagnetic emissions |
| --- |
| SonoEye is intended for use in the electromagnetic environment specified below. The customer or the user of SonoEye Series should assure that it is used in such an environment. |
| Emissions test | Compliance | Electromagnetic environment – guidance |
| RF emissions CISPR 11 | Group 1 | SonoEye uses RF energy only for its internal function. Therefore, its RF emissions are very low and are not likely to cause any interference in nearby electronic equipment. |
| RF emissions CISPR 11 | Class A | SonoEye is suitable for use in all professional healthcare environments, like physicians, offices, dental offices and those not directly connected to the public low-voltage power supply net work that supplies buildings used for professional medical purposes. |

| 2. Guidance and manufacturer’s declaration – electromagnetic immunity |
| --- |
| The SonoEye is intended for use in the electromagnetic environment the SonoEye should assure that it is used in such an environment. |
| Immunity test | IEC 60601 test level | Compliance level | Electromagnetic environment – guidance |
| Electrostatic discharge (ESD) IEC 61000-4-2 | ±8 kV contact ±2 kV, ±4 kV, ±8 kV, ±15kV air | ±8 kV contact ±2 kV, ±4 kV, ±8 kV, ±15kV air | Floors should be wood, concrete or ceramic tile. If floors are covered with synthetic material, the relative humidity should be at least 30 %. |
| Power frequency frequency (50/60 Hz) magnetic field IEC 61000-4-8 | 30 A/m | 30 A/m | Power frequency magnetic fields should be at levels characteristic of a typical location in a typical commercial or hospital environment. |
| Radiated RF EM fields IEC61000-4-3 | 3V/M 80MHz-2.7GHz 80%AM at 1KHz | 3V/M 80MHz-2.7GHz 80%AM at 1KHz | Radiated RF EM fields should be at levels characteristic of a typical location in a typical commercial or hospital environment. |
| Proximity fields from RF wireless communications equipment IEC 61000-4-3 | See teble 9 | See table 9 | Proximity fields from RF wireless communications equipment should be at levels characteristic of a typical location in a typical commercial or hospital environment. |

# Text boxes

4535 619 09131_A/795 * APR 2017

CAUTION: When the above life time is expired, the effectiveness and safety of system and transducers maybe greatly affected, so it’s NOT suggested to continue using the system and transducers even the system and transducers seem work properly. But if user still wants to continue using the system and transducers, user should first contact CHISON service center at CHISON headquarter to arrange the necessary safety check and calibration by CHISON’s authorized service engineer. If CHISON headquarter service center provides the calibration certificate for the related system or transducer, then user could continue use the system or transducers according to the calibration certificate. However, if CHISON headquarter service center concludes that the system or transducer is no longer complied to the safety and effectiveness standard, then user should immediately stop using the system or transducer. User understands that such check and calibration cost will be born by the user.

Systems and transducers keep on using after the life time may also be difficult to repair and maintain, so it’s suggested to renew the product after the life time.

Scanhead (Lens)

Transducer